Appendix-1: Methods and Quality Control (QC)

### A companion to:

Kjell Gundersen, Rocio Castano Primo, Tone Falkenhaug, Gayantonia Franze, Terje Jåvold, Linda F. Lunde, Arnfinn Morvik, Jane S. Møgster, Lars-Johan Naustvoll, Lena Omli, Ann-Kristin Olsen, Hilde A. Øyjordsbakken, Asuka Yamakawa (2025) Biogeochemical time-series and surveys in the North Sea, Skagerrak and Kattegat, 1990 – 2024

Background

The Plankton Chemistry Laboratory at the Institute of Marine Research (IMR) maintains quality control of precision and accuracy by daily assessments of analytical standard curves and the use of internal standards. As large data sets are becoming increasingly available for a wider, global array of users, it is important that the scientific community agree on a set of common practices and quality control (QC). Our laboratory find it crucial to maintain contacts with other laboratories through regional intercalibration studies such as QUASIMEME (http://www.quasimeme.org/) and in global intercalibration studies such as the IOCCP (https://repository.oceanbestpractices.org/handle/11329/883).

This dataset is a result of 30 years of work at sea and a great number of people have contributed to this product. We thank all onboard technicians and researchers that have collected biogeochemical samples over the years at IMR, including those who performed subsequent analysis at sea and in the land-based laboratories. The monitoring program in its current form was initiated by Francisco “Pancho” Rey and Lars Føyn. Together with Magnar Hagebø, they established most of the sample collection routines and analytical procedures at the Plankton Chemistry Laboratory in Bergen, that are still maintained to this date. A sizeable part of the time-series data from the North Sea, Skagerrak and Kattegat, were collected in the monitoring programs initiated by Didrik S. Danielssen at IMR Flødevigen Research Station, outside Arendal, and Lars Føyn at IMR in Bergen. The latest addition was the North Sea Ecosystem Survey (Økosystemtokt i Nordsjøen), established by Tone Falkenhaug at IMR in 2006.

The Environmental Survey in the North Sea (Miljøtoktet) started as a spring-bloom survey in 1988, with a focus on Skagerrak and Kattegat. At that time, a number of transects radiating from the coast of southern Norway, were also sampled late in the year. The North Sea Ecosystem Survey was established in 2006, as an April-May cruise, and the two surveys ran in tandem until 2014. From 2015, the two spring surveys merged and became the North Sea Ecosystem Survey, covering the northern half of the North Sea, Skagerrak and the northern part of Kattegat. Post-cruise reports, with a summary of sampling activities and immediate results, are issued from most of the surveys (see Selected cruise reports at the end of appendix). A selected number of transects from the North Sea Ecosystem Survey, are also covered 3-4 other times of the year, as an operational part of the Open Ocean Monitoring Programs (Snittundersøkelsene) at IMR.

Most of the surveys in the North Sea and adjacent regions have been funded by the Norwegian Ministry of Trade Industry and Fisheries (Nærings- og Fiskeridepartementet) and the Norwegian Environment Agency (Miljødirektoratet). The data set also contain nutrient samples collected from external grants that were analyzed by the Plankton Chemistry Laboratory and, as such, some data sets may also include contributions from the University of Bergen, University of Oslo, Norwegian Institute for Water Research.

Sample collection and QC

Annual cruise reports show the full magnitude and wide specter of samples collected on the North Sea surveys (see Selected cruise reports below), but only biogeochemical measurements are reported here. Seawater samples are collected from Niskin-type water bottles at predetermined depths, triggered by a CTD (Conductivity Temperature Depth) unit mounted on a rosette. Total number of stations sampled can vary between years and, in addition to random visits to the region on short-term projects, IMR has made repeated, multiple visits to the Torungen-Hirtshals transect (sampled 10-12 times a year). A selected number of transects, the Utsira-Start Point and the Hanstholm-Aberdeen transects, are also sampled several (3-4) times a year. These visits are additions to the annual, full-scale survey covering major parts of the North Sea, Skagerrak, and the northern part of Kattegat.

The Plankton Chemistry Laboratory is using a QC-flagging system to account for the quality of all the data generated (Jaccard et al. 2018) and by using  reference table 2 in OceanSITES Data Format Reference Manual (http://www.oceansites.org/docs/oceansites_data_format_reference_manual.pdf). Data with flags 1, 2 and 5 are made available for use in this publication, while flags 3 and 4 are deemed compromised and are not included. Due to the many different people used in the sampling program on IMR cruises, we are aware that minor mistakes can be made. Some of these can be corrected (e.g. a mislabeled depths with a value that clearly belongs somewhere else) and are given Flag = 5, while others are beyond correction (e.g. samples that shows signs of contamination or appears to have been stored too long) and are given flags 3 or 4. An important QC flag is number 2, where we label data points that are outside expected value, but for no apparent reason. Flag number 2 data can be an expression of small term changes that may turn out to be significant over larger time scales, and this kind of data cannot be discarded in ocean monitoring studies lasting decades to centuries. Unfortunately, data from samples collected prior to 2011 discarded flag number 2 labels along with flags 3 and 4. Therefore, flags number 2, 3 and 4 were routinely excluded from data sets during the 1990-2010 period.

Dissolved inorganic nutrients (Nitrite, Nitrate, Phosphate, Silicate)

After three rinses, each water sample (20 mL) is collected in a polyethylene vial. Up until the turn of the century (1999-2002) most nutrient samples were analyzed in real time onboard the ships and without chloroform additions. As the research cruise activity expanded, the number of automated analyzers could no longer match the number of ships operating simultaneously and nutrient samples were added 200 µL chloroform to subdue biological activity and stored dark at +4 ºC prior to analysis in the home laboratory within 1-6 weeks. Prior to analysis on an Automated Analysis (AA) system, the samples were acclimated to room temperature as the evaporating chloroform was evacuated by vacuum. At the Plankton Chemistry Laboratory in Bergen, a number of homemade Skalar and Alpkem hybrid systems have been used over the years (1990-2019). Initially, the Plankton Chemistry Laboratory at Flødevigen Research Station had a Chemlab-system (1989-2002), before the autoanalyzer was replaced with an Alpkem hybrid system. Colorimetric determinations of dissolved inorganic nutrients are based on the methods first described by Bendschneider, K. & Robinson, R.I. (1952) and Grasshof (1965) with minor adjustments suggested by the manufacturers (Alpkem, Chemlab, Skalar). Briefly, nitrate in seawater is reduced to nitrite coupled to a diazonium ion and, in the presence of aromatic amines, the resulting blue azo-dye is determined spectrophotometrically at 540 nm. The nitrate concentration is corrected for ambient nitrite (same analytical method as for nitrate, but without cadmium reduction) measured concurrently. Phosphate reacts with molybdate at low pH and the resulting phosphomolybdate is reduced with ascorbic acid to a blue complex measured spectrophotometrically at 810 nm (both laboratories). The Plankton Chemistry Laboratory in Bergen acquired a new Skalar AA system in 2017, where the phosphomolybdate peak is measured at 880 nm. Silicate (silicic acid) reacts with molybdate at low pH and the resulting silico-molybdate is reduced by ascorbic acid to a blue dye measured spectrophotometrically at 810 nm.

Dissolved ammonium samples (AMON)

After three rinses, each unfiltered water sample is collected in a 20 mL polyethylene vial (Bergen) or in a 25 mL brown glass flask (Flødevigen) and stored at -20 ºC. Prior to analysis, the sample is thawed to room temperature. At the Plankton Chemistry Laboratory in Bergen, dissolved ammonium (NH4+) is left to reacts with orthophtaldialdehyd (OPA) and the end-product (ammonium-OPA), is measured fluorometrically at 430 nm at 365 nm excitation using an Alpkem Flow Solution IV autoanalyzer (Kérouel & Aminot 1997, Holmes et al. 1999). At the Plankton Chemistry Laboratory in Flødevigen, the ammonium-indophenol end-product is measured spectrophotometrically at 640 nm (Solorzano 1969).

Chlorophyll-a and Phaeopigment samples (ChlA, Phaeo)

A standard volume (265 mL) is collected from each depth, collected on a 25 mm GFF filter and stored frozen (-20 ºC) until analysis in the land-based laboratory. Up until recently, pigment samples were transported home by one of the cruise participants, as hand-luggage in a cooler with frozen cooler-elements. These days, pigment samples are brought back in specially designed coolers, with an internal temperature recorder, that is rated for -20 ºC for a minimum of 3 days. In the laboratory, the samples are thawed in 90 % acetone, and stored at +4 ºC overnight before analysis on a Turner Design 10AU fluorometer. Phaeopigments (Phaeo) are measured separately from ChlA, in a second reading of the sample after adding 3 drops of a weak acid (5 % HCl). The fluorometer is standardized regularly using a solid standard with known fluorescence, and in accordance with Holm-Hansen & Riemann (1978) and the manufacturer (Turner Design 1992). Up until 2008 the drift in the light-source was monitored annually, but from 2009 the solid standard has been recorded every time the fluorometer is used. In the Flødevigen Laboratory, only 100 mL volumes are filter-collected and stored frozen (-20 ºC) until analysis, and only ChlA is measured from the acetone-extracted samples (Holm-Hansen et al. 1965, Jeffery & Humphrey 1975).

Dissolved oxygen samples by Winkler titrations (DOW)

Samples for dissolved oxygen are collected in dedicated glass BOD bottles (approximate volume 125 mL) and filled bottom-up using a Tygon-tubing. The sample is let to overflow approximately 3 times the volume and great care is taken to avoid small air bubbles inside of the sample bottle during filling. Thiosulfate titrations of dissolved oxygen are still done as first described by by Winkler (1888) but the method has seen some updates and improvements in later years (Carpenter 1965, Murray et al. 1968, Strickland & Parsons 1968). Grasshof et al. (1983) describe in detail the current method of sample collection, pretreatment and titrations of Winkler samples. Briefly, dissolved oxygen is reacting with an alkaline solution (Reagent 1) to form a manganese-hydroxy-complex. Under alkaline conditions, the Mn-complex is reacting with the iodide solution (Reagent 2) and left to precipitate in the sample bottle. The sample is added sulphuric acid (Bergen) or hydrochloric acid (Flødevigen) to dissolve the iodide precipitate (pH=1-2.5) and the yellow iodine is titrated by thiosulfate to a clear solution. The titrant is standardized by a known concentration of potassium iodate (KIO3) as described by Grasshof et al. (1983).

Literature

Carpenter, J.H. (1965) The Chesapeake Bay Institute. Technique for the Winkler oxygen method. Limnology & Oceanography 10:141-143

Bendschneider, K. & Robinson, R.I. (1952) A new spectrophotometric method for the determination of nitrite in seawater. Journal of Marine Research 2:87-96

Grasshoff, K. (1965) On the Automatic Determination of Phosphate, Silicate and Fluoride in Seawater. ICES Hydrographic Committee Report No. 129

Holm-Hansen, O. & Riemann, B. (1978) Chlorophyll a determination: improvements in methodology. Oikos 30:438-447

Holm-Hansen, O., Lorenzen, C.J., Holms, R.W. & Strickland, J.D.H. (1965) Fluorometric Determination of Chlorophyll. Journal du Conseil / Conseil Permanent International pour l'Exploration de la Mer 30: 3-15

Holmes, R.M., Aminot, A., Kérouel, R., Hooker, B.A., & Peterson, B.J. (1999) A simple and precise method for measuring ammonium in marine and freshwater ecosystems. Canadian Journal of Fisheries and Aquatic Science 56:1801–1808

Horne, R.A. (1969) Marine Chemistry, 568pp. Inter-science, New York, 1

Jaccard, P., Hjermann, D.Ø., Ruohola, J., Marty, S., Kristiansen, T., Sørensen, K., Kaitala, S., Mangin, A. & Poulinquen, S. (2018) Quality Information Document for global ocean reprocessed in-situ observations of biogeochemical products. Issue 2.0, 77 pp

Jeffery, S.W. & Humphrey, G.F. (1975) New spectrophotometric equations for determining chlorophylls a, b, c1 and c2 in higher plants, algae and natural phytoplankton. Biochemie und Physiologie der Planzen 167: 191-194

Kérouel, R. & Aminot, A. (1997) Fluorometric determination of ammonia in sea and estuarine waters by direct segmented flow analysis. Marine Chemistry 57:265–275.

Murray, J.N., Riley, J.P. & Wilson, T.R.S (1968) The solubility of oxygen in Winkler reagents used for the determination of dissolved oxygen. Deep-Sea Research 15:237-238

Solorzano (1969) Determination of ammonia in natural waters by the phenylhypochlorite method. Limnology & Oceanography 14:799-801

Strickland, J.D.H & Parsons, T.R. (1972) A Practical Handbook of Seawater Analysis. Fisheries Research Board of Canada, Bulletin 167:71-75

Turner Designs (1992) Model 10-AU-005 Field Fluorometer User’s Manual, Version S1C. Turner Designs, California, USA, 141 pp

Winkler, L.W. (1888) Die Bestimmung des im wasser gelösten Sauerstoffes. Berichte der Deutschen Chemishe Gesellschaft 21:2843-2855

Selected cruise reports

The Environmental Survey in the North Sea (Miljøtoktet) 2006:

https://www.hi.no/resources/publikasjoner/toktrapporter/2006/nr._7_2006.pdf

The North Sea Ecosystem Survey (Økosystemtokt i Nordsjøen) 2016:

https://www.hi.no/resources/publikasjoner/toktrapporter/2016/northsea_ecosystem_cruise_2016.pdf